## Slide 1: Critical Design review
Samuel Ellison
Group 8

## Slide 2: Project overview
Section 1

## Slide 3: Theory: Thermistors
Thermistors
Resistor with resistance dependent on temperature
Two types
  NTC: resistance decreases with increasing temp (this case)
  PTC: resistance increases with increasing temp
Notes: I’ll start by discussing the theory of thermistors. A thermistor is simply a resistor with a resistance that changes with the ambient temperature. The type of thermistor determines how the resistance changes with temperature. The two types are Negative and Positive Temperature Coefficient Thermistors. The relationships are shown in the graph here: resistance decreases with an increase in temp for NTC, and increases with an increase in temp for PTC.

## Slide 4: Theory: NTC T-R Models
Steinhart-Hart Equation
Third order approximation, constants can be solved with three R-T data points:
Parameter Equation
Alternative form of S-H with
value given
  1 ref point needed, nominal data and
value unknown (required calibration)
  2 ref point needed
Notes: The type of thermistor I will be using is an NTC. There are mainly two models that are used to calculate temperature from an NTC thermistor: SH and beta parameter. The SH is a third order approximation and uses three coefficients unique to the thermistor. The beta parameter is an alternative and simpler form of SH that uses just one coefficient: the beta parameter. It’s assumed in this model that the last coefficient c2 in SH is about 0, which is the case most of the time. The SH coefficients and beta parameter are sometimes given in the manufacturer specs. However, I will be deriving my coefficients from a calibration of the thermistor using three measured temperatures and resistances, which I will discuss in detail later on. All temperatures must be in Kelvin when using these models.

## Slide 5: Theory: Voltage Dividers
Voltage Divider
  Circuit producing output voltage proportional to input
  Used to measure resistance of unknown resistor
Thermistor Resistance
  Voltage divider used to calculate thermistor resistance
Equation
If is the unknown thermistor resistance
measured by Arduino, V, is resistance of known resistor
Notes: Voltage Dividers are extremely useful circuit components, especially when working with Arduino. The circuit produces an output voltage proportional to the input voltage based on resistance values R1 and R2. The purpose of using a voltage divider in general is to measure the resistance of an unknown resistor (R2 in circuit). When applied to Arduino, voltage dividers are important because the digital and analog signals going to the microcontroller can only read voltage output, not resistance. So measuring the thermistor’s resistance requires the use of a voltage divider. GO THROUGH EQUATIONS. Voltage dividers in general do not have to use resistors, but in this project they do.

## Slide 6: Theory: Thermal Time Constant
Thermal time constant is time required for thermistor to reach 63.2% of difference between initial and final ambient temperatures
Assuming a lumped capacitance model for transient temperature analysis,
where and are the ambient temperatures and is the thermistor’s reading at time .
If , then

## Slide 7: Theory: Steady-State Condition
Temperature readings can be taken when steady state condition is reached
Condition is satisfied when derivative is low enough
Exact derivative not possible, need for accurate approximation
Five-Point Stencil for First Derivative
  Central Difference, Fourth Order Accuracy
  If 5 most recent temps are , the derivative at time is about
  is the time step, equal to where is sampling rate
  Derived From Taylor Series

## Slide 8: Requirements
Objective: Design an electro-thermal system using Arduino kit with NTC thermistor
Accurate temperature measurements at ice, steam, and room temp water
Thermistor extension from microcontroller module
  Ease of application to surface
  User and Thermistor Protection
Both analog and digital signals must be used
User input for
  Sampling rate, error message if limit is exceeded
Digital temperature display in C and F
Error must stay within a certain tolerance (5%)
User must be able to start/stop sampling
User must be alerted when steady state temperature is reached
Notes: The overall objective of this project is to design a thermometer from an Arduino kit with an NTC thermistor. The requirements of this project are (go through points)

## Slide 9: Concept of Operations
Sampling Rate Specification
  User prompted input in Serial dialogue (Hz)
  While loop to delay void loop by correct time to ensure given rate
  Error message if invalid (more than limit)
  Limit determination
    Time one program loop : (micros())
    Sampling rate limit : (conversion)
Start/Stop Buttons
  Start button begins/resumes data collection, green LED on
  Stop button terminates data collection, green LED off
Thermistor Placement
  User to press thermistor against solid or dip into fluid
  Need not to hold thermistor directly to avoid injury
Liquid Crystal Display (LCD)
  Potentiometer to vary contrast
  Celsius and Fahrenheit displayed
Steady State Identification
  Blue LED Indicator
  Most recent 5 temperatures stored in array
  Criterion
    First derivative calculation with 5-point stencil

## Slide 10: Expected Results
Calibration 1
  Nominal Resistance
  T-R data
Time Constant
  Should be s (manufacturer specs)
Temperature error
  Percent error around 5%
  Expect something higher
Sampling Rate Limit
  Loop Run Time 10 ms
  100 samples/sec
Point | Resistance () | Temperature ()
1 | 27101.45 | 3.1
2 | 11157.02 | 21.9
3 | 1390.16 | 83.2
Point |  | 
1 | 27101.45 | 3.1
2 | 11157.02 | 21.9
3 | 1390.16 | 83.2
Notes: Note c2 is small compared to others, beta parameter model assumption valid.

## Slide 11: System overview
Section 2

## Slide 12: Requirement/Design Changes
Start/Stop more accurate description than On/Off for buttons
While loop instead of delay command
  Accuracy
  micros() rather than millis()
Changed thermistors
  Necessary recalibration
S-H and Parameter Accuracy/Error Comparison
Steady State condition automatically stops measurements
  Used for calibration
De-scopes
  Consolidation of Code
    Decreases minimum run time, increases efficiency
  Reorganization of wiring
    Simpler/Easier on the eyes

## Slide 13: Mechanical Design Elements
Extension from Breadboard
  Dupont Wires connected to breadboard
  Four total wires used, two for each terminal
Rigidity/Flexibility
  Bottom wires left flexible
  Top wires made rigid with coffee straws
  Free thermistor movement, hold at bottom of straws
Protection/Adhesives
  Electrical Tape wrapped around top wires
  Nail polish

## Slide 14: Electrical Design Elements
Power Supply
  3.3 V is supplied to breadboard
  AREF used to determine accurate voltage supplied
LCD
  Temperatures calculated in program and sent to LCD for user
  Uses Digital ports 2,3,4,5,11,12 to send information to LCD
  Potentiometer controls contrast, Resistor sets brightness
Start/Stop Buttons
  Program checks which buttons are sending signals to digital, uses port 6
  Buttons connected with resistors from ground to power
  Control green LED
LEDs
  Green LED: Power supplied when start button sends a signal (conditional)
  Blue LED: Power supplied when steady state condition is met
  Both digital signals
Voltage Divider and Thermistor
  Thermistor acts as “unknown” resistor in voltage divider circuit
  Output voltage sent to analog A0

## Slide 15: System Block Diagram
Microcontroller
Temperature Display (LCD)
On/Off Buttons
Thermistor Voltage Divider
Steady State Criterion
Arduino Program
Complete System
Potentio-
meter
Green LED
Blue LED
Rigid Extension & Protection
Notes: Go over color code: black: basic circuit components: pot, LED. Yellow: Subsystems of the design.
Go from left to right
Arduino program and microcontroller are sending information back and forth

## Slide 16: System Setup
Stop Button
Start Button
Green LED
Blue LED
LCD
Potentiometer
Thermistor
Extension
Notes: On/Off buttons: Send signals to start/end data collection
Potentiometer: Sends resistance information to affect contrast of LCD
Green LED: Turns on when system is running, i.e. on button has most recently been pressed. Turns off when system not running (off button)
LCD: Begins with blank screen. Contrast may be adjusted before on button. Once system is running, temperatures are displayed. When system is off, temperatures disappear, blank screen.
Thermistor: Sends information through voltage divider to program which converts VRT. Calculations and displays on LCD only made when system is ON
Blue LED: Can only work when system is ON. Turns on when SS reached, which is condition defined in program

## Slide 17: Subsystem design
Section 3

## Slide 18: Temperature Display (LCD)
Output temperature reading to LCD
Digital Signal
Given in Fahrenheit and Celsius
Signal sent when
  Sampling rate defined
  Start button pressed
Variable contrast
Constant brightness
TEMP: ___
TEMP: ___
POWER
Potentiometer
Resistor
CONTRAST
BRIGHTNESS
TEMPERATURE CALCULATION
LCD
GROUND
DIGITAL SIGNAL PROCESSING

## Slide 19: Start/Stop Buttons
Start button
  Begins loop calculating temps which sends to LCD
  Turns green LED on
Stop button
  Terminates program and thermistor readings
  Turns green LED off
Pressing button sends signal to microcontroller for limited time
Each loop checks which buttons send signal
DETAILS:
  Start button immediately begins sampling
  Stop button must be held until light turns off
    Low sampling rates
POWER
GROUND
10
Buttons
DIGITAL SIGNAL PROCESSING
LED
STOP
START
10
220
CONDITIONALS

## Slide 20: Thermistor Voltage Divider
Produces Output Voltage dependent on thermistor resistance
Arduino program to calculate thermistor resistance
Uses 10kOhm resistor in series with thermistor
3.3V input voltage
  Exact voltage measured with AREF EXTERNAL
  3.3 V vs 5 V
    5 V input fluctuates more based on:
      Length of USB
      Level of battery charge
    3.3 V highly conditioned by hardware
    AREF reduces error
Analog signal A0
10
THERMISTOR
POWER
GROUND
ANALOG SIGNAL PROCESSING
TEMPERATURE CALCULATION

## Slide 21: Steady State Criterion
Program stores 5 most recent temperatures
Five-point stencil derivative calculation
Derivative meets SS criterion, signal sent
Blue LED indicates SS
Calibration
  Break out of void loop when SS is reached so SS temperature stays on LCD
LED
DIGITAL SIGNAL PROCESSING
GROUND
220
DERIVATIVE
CALCULATION

## Slide 22: Rigid Extension/Protection
Extension
  Dupont Wires to extend thermistor from breadboard
    Can use multiple connections to extend further distance
Protection
  Coffee stirring straws surrounding ribbon wires = rigidity
  Nail polish coat on thermistor wires
  Electrical tape to surround wiring/connections
    Extra layer
BREADBOARD
STRAWS
FLEXIBLE
RIGID
TAPE

## Slide 23: Prototyping and analysis
Section 4

## Slide 24: Prototyping Results: LCD
Displays temperatures in C and F
Potentiometer effectively controls contrast
Does not start showing temperatures until start button is pressed

## Slide 25: Prototyping Results: LEDs
Start/Stop LED
  Pressing Start turns green LED on
  Pressing Stop turn green LED off
Steady-State LED
  If five-point stencil derivative drops below 0.001, blue light turns on
Continuous Sampling
  Sampling continues even after SS reached
  Only way to terminate is with stop button
Non-continuous Sampling
  Sampling terminates when SS reached
  “Grace period” of 5 seconds once started
  Stop button can also terminate sampling before SS

## Slide 26: Prototyping Results: Buttons
Buttons not operable until sampling rate inputted
Start button begins sampling and temp display on LCD and turns on green LED
Stop button ends sampling, freezes temp display on LCD, and turns off green LED
Start button immediately begins sampling
Stop button should be held down to ensure termination of sampling

## Slide 27: Prototyping Results: Voltage Divider
Nominal Resistance around 10kOhm
  Use 10kOhm resistor in series
Effectively converts thermistor resistance to a readable voltage

## Slide 28: Analysis Results: Calibration
Point | Resistance () | Temperature ()
1 | 26967.51 | 1.6
2 | 10197.24 | 24.5
3 | 1022.6 | 94.7
Point |  | 
1 | 26967.51 | 1.6
2 | 10197.24 | 24.5
3 | 1022.6 | 94.7
Reference Points
Calibration Coefficients
Coefficient | Value
 | 8.9605e-04
 | 2.5788e-04
 | 1.0602e-07
 | 3.4730e+03
Coefficient | Value
 | 8.9605e-04
 | 2.5788e-04
 | 1.0602e-07
 | 3.4730e+03
model
S-H model

## Slide 29: Analysis Results: Test Data
Point | Resistance () | () | () | ()
Ice | 27372.26 | 1.27 | 1.28 | 0.9
Room | 10562.25 | 23.61 | 23.61 | 23.6
Steam | 905.22 | 99.26 | 102.46 | 98.2
Point |  |  |  | 
Ice | 27372.26 | 1.27 | 1.28 | 0.9
Room | 10562.25 | 23.61 | 23.61 | 23.6
Steam | 905.22 | 99.26 | 102.46 | 98.2
1.079% diff
4.338% diff
Models Deviate at Higher Temperatures
  Beta model calibrated using low temps
  Hence high percent diff at high temps

## Slide 30: Analysis Results: Sampling Rate Limit & Time Constant
Sampling Rate Limit
Time for a single loop
  microseconds
Maximum sampling rate
Set limit to (margin of error)
Time Constant
Trial | 
1 | 10.39s
2 | 8.76s
3 | 9.42s
Avg | s
Trial | 
1 | 10.39s
2 | 8.76s
3 | 9.42s
Avg | 
Trial | 
1 | 10.48s
2 | 9.73s
3 | 9.92s
Avg | s
Trial | 
1 | 10.48s
2 | 9.73s
3 | 9.92s
Avg | 
Cold to Ambient
Hot to Ambient

## Slide 31: Error Analysis: Temperature
Propagation of Uncertainty
  Thermistor Resistance from Voltage Divider Equation
    Given uncertainties
      Known Resistor:
      Voltages:
  S-H coefficients
    Ref. Thermometer Temps:
    Thermistor Resistances at 3 Calibration Points
  parameter
    Ref. Thermometer Temps:
    Thermistor Resistances at 2 Calibration Points
  Temperature Measurements
    S-H
      Coefficients c0, c1, c2
      Thermistor Resistance
      Coefficient
      Reference Temp
      Reference Resistance
      Thermistor Resistance
Point | Temp Error (SH) | Temp Error ()
Ice | 10.078% | 0.220%
Room | 10.322% | 0.135%
Steam | 11.471% | 0.635%
Point | Temp Error (SH) | 
Ice | 10.078% | 0.220%
Room | 10.322% | 0.135%
Steam | 11.471% | 0.635%
Higher Error in SH from more propagation
Beta less accurate at higher range
Beta accurate at temperatures near calibration points

## Slide 32: Conclusions & future testing
Section 5

## Slide 33: System Level Testing
Conclusions
  S-H model best for this system
    Gives more accurate temperatures over required range (ice to steam point) than beta model
    Percent Error is very high
      Need near perfect reference thermometer to prevent heavy propagation of error in coefficients
  Ideal sampling rate around 3-4 Hz
    Higher rates make blue LED flash off and on too frequently
    Lower rates don’t show enough intermediate temperatures
  Time constant and beta parameter calibration consistent with company specs
Entire System performs requisite function
  Accurate temperatures are measured and displayed at ice and steam points of water with S-H model
  User able specify sampling rate easily
  Thermistor held by rigid fixture
  Taking measurements is safe for the user and equipment

## Slide 34: Mechanical Elements
Dupont Wires
  Requirements:
    Thermistor Extension from Breadboard
Thin Coffee Straws
  Requirements:
    Easy thermistor movement and application
    Wire protection
Electrical Tape
  Requirements:
    Thermistor/Wire/User protection
    Ensures secure connection
Nail Polish
  Requirement:
    Thermistor Wire protection
Future Testing
  Clamp/Clip to edge of beaker/pot/bowl/cup
    Hands free
  Battery power supply
    System mobility
  Microcontroller/breadboard attachment to same surface
    Ease of transportation of device
  Thermistor Casing
    Protection when not in use

## Slide 35: Electrical Elements
LCD
  Requirements:
    Digital signal used
    Digital temperature display in C and F
Start/Stop Buttons
  Requirements:
    Digital Signal used
    User can start/stop sampling
Green LED
  Requirements:
    Digital Signal used
    Indication that sampling is occurring
Blue LED
  Requirements:
    Digital Signal used
    Indication that SS has been reached
Voltage Divider
  Requirements:
    Analog Signal used
    Digital Temperature display
Future testing
  Sampling rate input on LCD
  On/Off Buttons to completely turn on/off LCD

## Slide 36: Software Elements
Serial Monitor
  Requirements:
    User input for sampling rate
Derivative Calculation
  Requirements:
    Indication that SS has been reached
Conditional Statements
  Requirements:
    User can start/stop sampling
    Indication that sampling is occurring
    Indication that SS has been reached
    Error message if sampling rate limit exceeded
While loop delay
  Requirements:
    User input for sampling rate

## Slide 37: Questions?

## Slide 38: Video Links
https://youtu.be/dhY4Hbpf9sk
https://youtu.be/hHX3oRhQVSM
https://youtu.be/eewrN2xHJlc

## Slide 39: Appendix

## Slide 40: Preliminary Design review
Samuel Ellison
Group 8

## Slide 41: Project overview
Section 1

## Slide 42: Theory: Thermistors
Thermistors
Resistor with resistance dependent on temperature
Two types
  NTC: resistance decreases with increasing temp (this case)
  PTC: resistance increases with increasing temp
Notes: I’ll start by discussing the theory of thermistors. A thermistor is simply a resistor with a resistance that changes with the ambient temperature. The type of thermistor determines how the resistance changes with temperature. The two types are Negative and Positive Temperature Coefficient Thermistors. The relationships are shown in the graph here: resistance decreases with an increase in temp for NTC, and increases with an increase in temp for PTC.

## Slide 43: Theory: NTC T-R Models
Steinhart-Hart Equation
Third order approximation, constants can be solved with three R-T data points:
Parameter Equation
Alternative form of S-H with
Requires only one R-T data point
usually very small, accurate proxy
value given in manufacturer specs
Notes: The type of thermistor I will be using is an NTC. There are mainly two models that are used to calculate temperature from an NTC thermistor: SH and beta parameter. The SH is a third order approximation and uses three coefficients unique to the thermistor. The beta parameter is an alternative and simpler form of SH that uses just one coefficient: the beta parameter. It’s assumed in this model that the last coefficient c2 in SH is about 0, which is the case most of the time. The SH coefficients and beta parameter are sometimes given in the manufacturer specs. However, I will be deriving my coefficients from a calibration of the thermistor using three measured temperatures and resistances, which I will discuss in detail later on. All temperatures must be in Kelvin when using these models.

## Slide 44: Theory: Voltage Dividers
Voltage Divider
  Circuit producing output voltage proportional to input
  Used to measure resistance of unknown resistor
Thermistor Resistance
  Voltage divider used to calculate thermistor resistance
Equation
If is the unknown thermistor resistance
measured by Arduino, V, is resistance of known resistor
Notes: Voltage Dividers are extremely useful circuit components, especially when working with Arduino. The circuit produces an output voltage proportional to the input voltage based on resistance values R1 and R2. The purpose of using a voltage divider in general is to measure the resistance of an unknown resistor (R2 in circuit). When applied to Arduino, voltage dividers are important because the digital and analog signals going to the microcontroller can only read voltage output, not resistance. So measuring the thermistor’s resistance requires the use of a voltage divider. GO THROUGH EQUATIONS. Voltage dividers in general do not have to use resistors, but in this project they do.

## Slide 45: Theory: Thermal Time Constant
Thermal time constant is time required for thermistor to reach 63.2% of difference between initial and final ambient temperatures
Assuming a lumped capacitance model for transient temperature analysis,
where and are the ambient temperatures and is the thermistor’s reading at time .
If , then

## Slide 46: Requirements
Objective: Design an electro-thermal system using Arduino kit with NTC thermistor
Thermometer measurement of freezing/boiling water (STP)
Thermistor extension from microcontroller module (10cm)
  Rigid fixture, allowing for ease of application to touch solid surface or penetrate fluid
  Thermistor protection from water damage
Both analog and digital signals must be used
User input for
  Sampling rate, error message if limit is exceeded
Digital temperature display in C and F
Error must stay within a certain tolerance (TBD)
User must be able to turn on/off the device
Notes: The overall objective of this project is to design a thermometer from an Arduino kit with an NTC thermistor. The requirements of this project are (go through points)

## Slide 47: Concept of Operations
Sampling Rate Specification
  User prompted input in Serial dialogue (Hz)
  Delay each loop by 1/input seconds
  Error message if invalid (more than limit)
  Limit determination
    Time one program loop : (millis command)
    Sampling rate limit : (conversion)
On/Off Buttons
  On button begins data collection, green LED on
  Off button terminates data collection, green LED off
Thermistor Placement
  User to press thermistor against solid or dip into fluid
  Need not to hold thermistor directly to avoid injury
Liquid Crystal Display (LCD)
  Potentiometer to vary contrast
  Celsius and Fahrenheit displayed
Steady State Indication/Criterion
  Blue LED
  Most recent 5(+) temperatures stored in array
  SS when all temps within 0.1 of another

## Slide 48: Operations/Tasks for Designer
Thermistor Calibration
  Voltage Divider to measure thermistor nominal resistance  choose second resistor
  Record thermistor resistance at three (freezing, room, boiling) known temps
    Use reference thermometer to record temps
  Three data points to calculate S-H constants (room temp for parameter)
Error Propagation and Uncertainty Analysis
  Factors in calculation of Temperature
    Thermistor Resistance
      Resistor 1 resistance
      Input Voltage
      Output Voltage
    S-H coefficients
      Three temperatures from reference thermometer
      Three thermistor resistances
  Different errors for C and F
Time Constant Measurement
  Take room temp reading
  Bring temp up to about 100C (boiling water)
  Remove thermistor from water, start timer
  Record time when temp = 63.2% diff
  Programmable
  Error Propagation

## Slide 49: Expected Results
Calibration
  Nominal Resistance
  T-R data
Time Constant
  Should be s (manufacturer specs)
Temperature error
  ideal (SS criterion)
  Expect something higher
Sampling Rate Limit
  Time to run one loop of program
  Possibly
  100,000 samples/sec
Point | Resistance () | Temperature ()
1 | 27101.45 | 3.1
2 | 11157.02 | 21.9
3 | 1390.16 | 83.2
Point |  | 
1 | 27101.45 | 3.1
2 | 11157.02 | 21.9
3 | 1390.16 | 83.2
Notes: Note c2 is small compared to others, beta parameter model assumption valid.

## Slide 50: System overview
Section 2

## Slide 51: Subsystem Definitions
Temperature Display
  How users will see/be able to read temperatures
On/Off buttons
  How users will turn display/program on and off
Voltage Divider
  How the thermistor is set up on the breadboard to determine its resistance
Thermistor Extension/Rigid Fixture
  How the thermistor will extend from module and its protection from environment
Steady State Criterion/Indicator
  How program will determine if steady state has been achieved and how users can determine if the temperature reading is stable and can be recorded
Arduino Code
  How the program is written to accomplish all objectives/requirements

## Slide 52: System Block Diagram
Microcontroller
Temperature Display (LCD)
On/Off Buttons
Thermistor Voltage Divider
Steady State Criterion
Arduino Program
Complete System
Potentio-
meter
Green LED
Blue LED
Rigid Extension & Protection
Notes: Go over color code: black: basic circuit components: pot, LED. Yellow: Subsystems of the design.
Go from left to right
Arduino program and microcontroller are sending information back and forth

## Slide 53: Critical Interfaces
Timing | System Component | Direction | Affects/Affected by
a | on/off buttons | input | begins/ends everything
a | Potentiometer | input | LCD
b | Green LED | output | on/off buttons
b | LCD | output | Potentiometer, on/off buttons
b | Thermistor | input//output | LCD, Blue LED // on/off buttons
c | Blue LED | output | on/off buttons, data collection rate
User must first input the sampling rate in serial monitor  defines delay in each loop
On/Off buttons: Send signals to start/end data collection
Potentiometer: Sends resistance information to affect contrast of LCD
Green LED: Turns on when system is running, i.e. on button has most recently been pressed. Turns off when system not running (off button)
LCD: Begins with blank screen. Contrast may be adjusted before on button. Once system is running, temperatures are displayed. When system is off, temperatures disappear, blank screen.
Thermistor: Sends information through voltage divider to program which converts VRT. Calculations and displays on LCD only made when system is ON
Blue LED: Can only work when system is ON. Turns on when SS reached, which is condition defined in program
Notes: I will walk through a sort of “typical” or “expected” cycle of temperature data collection using this chart and the descriptions below.

## Slide 54: System Setup
On Button
Off Button
Green LED
Blue LED
LCD
Potentiometer
Thermistor
Notes: On/Off buttons: Send signals to start/end data collection
Potentiometer: Sends resistance information to affect contrast of LCD
Green LED: Turns on when system is running, i.e. on button has most recently been pressed. Turns off when system not running (off button)
LCD: Begins with blank screen. Contrast may be adjusted before on button. Once system is running, temperatures are displayed. When system is off, temperatures disappear, blank screen.
Thermistor: Sends information through voltage divider to program which converts VRT. Calculations and displays on LCD only made when system is ON
Blue LED: Can only work when system is ON. Turns on when SS reached, which is condition defined in program

## Slide 55: Subsystem design
Section 3

## Slide 56: Temperature Display
Output temperature reading to LCD
Digital Signal
Given in Fahrenheit and Celsius
Requires “on” button
Variable contrast
Constant brightness
Risk:
  Using a potentiometer requires user to turn very small knob.
Mitigation:
  If this were to be a selling product, have versions with contrast already set with a resistor
TEMP: ___
TEMP: ___
POWER
Potentiometer
Resistor
CONTRAST
BRIGHTNESS
TEMP CALCULATION
TEXT
LCD
GROUND

## Slide 57: On/Off Buttons
On button
  Starts program and thermistor readings
  Turns green LED on
Off button
  Terminates program and thermistor readings
  Turns green LED off
  Clears LCD
Pressing button sends signal to microcontroller for limited time
Each loop checks which buttons send signal
Risk:
  Both buttons sending signals could cause problems
Mitigation:
  Create a condition to do nothing once one of the buttons turns off
POWER
GROUND
ON
OFF
Resistors
Buttons
DIGITAL SIGNAL PROCESSING
LED

## Slide 58: Thermistor Voltage Divider
Produces Output Voltage dependent on thermistor resistance
Arduino program to calculate thermistor resistance
Uses 10kOhm resistor in series with thermistor
5V input voltage
Analog signal
Risk:
  Using a resistor with resistance not close to nominal thermistor resistance may result in poor outputs/damage to thermistor
Mitigation:
  Use simple Voltage Divider to determine nominal resistance of thermistor first, then choose resistor accordingly before performing experiments at elevated/lowered temperatures
Resistor
THERMISTOR
POWER
GROUND
ANALOG SIGNAL PROCESSING

## Slide 59: Steady State Criterion
Program stores 5 most recent temperatures
Nested for loop to determine variance in values
If all values are within 0.1 Celsius, condition is met
Blue LED indicates SS
Risk:
  Sampling rate will change how often steady state is checked
Mitigation:
  Specify the sampling rate for steady-state check to be constant regardless of input.
LED
Resistor
DIGITAL SIGNAL PROCESSING
GROUND

## Slide 60: Rigid Extension/Protection
Extension
  Use Ribbon Wires to extend thermistor from breadboard
    Can use multiple connections to extend further distance
Protection
  Coffee stirring straws surrounding ribbon wires = rigidity
  Leave a bit of space before thermistor for rotational/dipping ability
  Nail polish to secure connections/extra layer of protection
  Electrical tape to surround wiring/connections
    Extra layer
Risk:
  Measuring extremely hot/cold bodies can be dangerous to user
Mitigation:
  Rigid extension allows user to hold bottom of wires far from sensor and still accurately position the thermistor

## Slide 61: Prototyping plan
Section 4

## Slide 62: Liquid Crystal Display
The LCD uses digital signal from microcontroller
Begin with basic setup and code
  Constant brightness/contrast using resistors
  Use basic commands in code to send messages to LCD
Shift to more variable interface
  Potentiometer to change contrast/brightness
Need to display calculated temperatures at specific locations on LCD

## Slide 63: LEDs
LED has two wires
Can be turned on/off when specific conditions are met
One end connects to resistor to ground
Other end connects to something to supply current
  Connect to voltage directly
  Connect to digital/analog signal
    When condition is met, signal is from microcontroller to LED, turns on
    Can also specify the LED to turn off with conditionals
ON/OFF LED
  Program conditionals to turn green LED on when ON button sends signal and OFF button is not sending signal
  Program conditionals to turn green LED off when ON button is not sending signal and OFF button sends signal
Steady-State LED
  Program a conditional to turn on blue LED when condition is met, turn off if not met

## Slide 64: Buttons
Buttons send a signal for limited time when pressed
Begin with experimenting with a single button to turn LED on
Program to ensure button stays on forever once button is pressed
Add another button to turn off LED when pressed.
Experiment with conditionals in code to find solution
Add LCD and data collection into loop
  Start data collection only if on button is pressed
  Stop data collection when off button is pressed
  Clear display when turned off
  Tell user to push on button to begin data collection

## Slide 65: Voltage Divider
Create a Voltage Divider with 2 constant resistors to find “unknown” resistance R2
Replace R2 with thermistor to determine nominal resistance
Replace R1 with appropriate resistor depending on nominal resistance
Add Voltage Divider with thermistor to entire circuit

## Slide 66: Project management
Section 5

## Slide 67: Schedule

## Slide 68: Budget
Costs
  Arduino Kit: $39.09
  Reference thermometer: $4.99
  Electrical tape: $9.99
  Straws: $4.99
  Total: $59.06

## Slide 69: Questions?